## PDPROJ-2-DesignInputs
ID | Title | UN ref | Description | Notes | Parent Component | Last Updated | Latest Design Review | Obsolete | Reason
PDPROJ-2 | RSD Guardian 0.9x
DI-1 | The device shall incorporate a sensor board with an RF sensor (FMCW) for detection of inmate motion and falls | UN-1 |  |  | RSD Guardian 0.9x | 2024-08-13T14:18:47.452Z |  | FALSE
DI-2 | The device shall incorporate an infrared camera to perform ambient temperature monitoring and possibly support movement detection | UN-1 |  |  | RSD Guardian 0.9x | 2024-08-13T14:19:04.868Z |  | FALSE
DI-3 | The devices sensor shall operate within 24 GHz, 60 GHz, 74 GHz and 120GHz frequency range. | UN-1 |  |  | RSD Guardian 0.9x | 2024-08-13T14:19:26.094Z |  | FALSE
DI-4 | The device's sensors shall be functional within at least 13 feet to cover 10x10x10 cell | UN-1 |  |  | RSD Guardian 0.9x | 2024-08-13T14:19:39.591Z |  | FALSE
DI-5 | The devices sensors shall be powered by power over ethernet (POE) requiring no more than 10 Watts power consumption (currently 7.5W, POE =52W). | UN-1 |  |  | RSD Guardian 0.9x | 2024-08-13T14:23:02.688Z |  | FALSE
DI-6 | The device shall incorporate a flash memory to store sensor data with the following minimum requirements: 128 Kb | UN-1 |  |  | RSD Guardian 0.9x | 2024-08-13T14:23:14.286Z |  | FALSE
DI-7 | The device shall incorporate a microprocessor\controller for the purpose of controlling the RF frontends and IR camera | UN-1 |  |  | RSD Guardian 0.9x | 2024-08-13T14:25:04.155Z |  | FALSE
DI-8 | The microprocessor shall have the following minimum requirements: TBD | UN-1 |  |  | RSD Guardian 0.9x | 2024-08-13T14:25:30.551Z |  | FALSE
DI-9 | The system shall send data and video over an ethernet cable for display by a cloud front end | UN-3 |  |  | RSD Guardian 0.9x | 2024-08-13T14:56:22.303Z |  | FALSE
DI-10 | The sensor shall be affixed within an enclosure no larger than the following dimensions: 10" x 10" x 3" | UN-4 |  |  | RSD Guardian 0.9x | 2024-08-13T14:56:48.499Z |  | FALSE
DI-11 | The device shall not weigh more than 2 pounds. | UN-4 |  |  | RSD Guardian 0.9x | 2024-08-13T14:56:57.713Z |  | FALSE
DI-12 | The enclosure shall be of the following material: polycarbonate | UN-4 |  |  | RSD Guardian 0.9x | 2024-08-13T14:57:09.022Z |  | FALSE
DI-13 | The enclosure shall include ingress protection in accordance with IP66 requirements and fire rating of UL94 | UN-4 |  |  | RSD Guardian 0.9x | 2024-08-13T14:57:20.299Z |  | FALSE
DI-14 | The enclosure shall connect to the POE on the top of the system near the ceiling mount to hide the connection | UN-4 |  |  | RSD Guardian 0.9x | 2024-08-13T14:57:30.500Z |  | FALSE
DI-15 | The mounting system shall be able to support the full weight of the device. | UN-5 |  |  | RSD Guardian 0.9x | 2024-08-13T14:57:58.011Z |  | FALSE
DI-16 | The device shall be able to withstand forces and impacts up to .5 Kg ball drop per IEC 60601-1. | UN-5 |  |  | RSD Guardian 0.9x | 2024-08-13T14:58:08.383Z |  | FALSE
DI-17 | Device will be operable under the environmental conditions below: Operating conditions: Temperature Range of +5 to +40 degrees C, relative humidity range of 15%-90%, and atmospheric pressure range of 700 hPa to 1060 hPa. | UN-6 |  |  | RSD Guardian 0.9x | 2024-08-13T14:58:40.491Z |  | FALSE
DI-18 | The device shall display an alert when it detects the absence of a prisoner in their cell | UN-7 |  |  | RSD Guardian 0.9x | 2024-08-13T14:59:08.010Z |  | FALSE
DI-19 | The display shall display an alert when the device detects a fall | UN-7 |  |  | RSD Guardian 0.9x | 2024-08-13T14:59:17.687Z |  | FALSE
DI-20 | The alerts shall be displayed on the Dashboard | UN-7 |  |  | RSD Guardian 0.9x | 2024-08-13T14:59:26.832Z |  | FALSE
DI-21 | Device labeling is compliant with FDA 21 CFR 801 and 21 CFR Part 1010 and includes the following information: Manufacturer Name and Address, Device Model, Description, Lot/Serial Number, Date of Manufacture, UDI, Rx Use Only Statement, and the electrical certification symbol. | UN-8 |  |  | RSD Guardian 0.9x | 2024-08-13T15:00:31.929Z |  | FALSE
DI-22 | Device labeling is compliant with IEC 60601-1, IEC 60601-1-2, and IEC 60601-1-11 and includes the following information: IP21 rating symbol, IEC 60601-1 Table D.1 Symbol 11 â€œConsult Accompanying Documentsâ€ Warning Symbol, the Rated Supply DC Voltage/ Amperage, the Rated Supply Input Voltage/Frequency/ Amperage. | UN-8 |  |  | RSD Guardian 0.9x | 2024-08-13T15:01:05.532Z |  | FALSE
DI-23 | The device labeling/IFU shall incorporate appropriate information on risks/hazards/ and mitigations as identified by the risk assessment. | UN-8 |  |  | RSD Guardian 0.9x | 2024-08-13T15:01:25.322Z |  | FALSE
DI-24 | Power shall be provided to the device by POE | UN-9 |  |  | RSD Guardian 0.9x | 2024-08-13T15:01:42.151Z |  | FALSE
DI-25 | The power adapter shall be compatible with Type B sockets (120 V, 60 Hz). | UN-9 |  |  | RSD Guardian 0.9x | 2024-08-13T15:01:53.556Z |  | FALSE
DI-26 | The power adapter shall convert the power supply from the power socket to an operating voltage of at 5V DC @ 4 A | UN-9 |  |  | RSD Guardian 0.9x | 2024-08-13T15:02:10.828Z |  | FALSE
DI-27 | The power adapter shall deliver power to the device via a USB-C cable. | UN-9 |  |  | RSD Guardian 0.9x | 2024-08-13T15:02:27.508Z |  | FALSE
DI-28 | The device shall have a use life of 1 year. | UN-10 |  |  | RSD Guardian 0.9x | 2024-08-13T15:03:37.375Z |  | FALSE
DI-29 | Device will be packaged so that device and included accessories, as applicable, are undamaged at point of use. | UN-11 |  |  | RSD Guardian 0.9x | 2024-08-13T15:04:12.612Z |  | FALSE
DI-30 | The device shall be packaged within a Cardboard Box + Foam, alongside the power adapter, power cable, and instructions for use. | UN-11 |  |  | RSD Guardian 0.9x | 2024-08-13T15:04:29.223Z |  | FALSE
DI-31 | The device packaging shall include an outer device label - See UN-4 unobtrusive | UN-11 |  |  | RSD Guardian 0.9x | 2024-08-13T15:04:40.847Z |  | FALSE
DI-32 | The finished device packaging shall be of the following dimensions and weight: TBD - see DI 10? | UN-11 |  |  | RSD Guardian 0.9x | 2024-08-13T15:04:54.360Z |  | FALSE
DI-33 | Device design will be compliant to IEC 60601-1 to ensure medical device electrical equipment safety | UN-12 |  |  | RSD Guardian 0.9x | 2024-08-13T15:05:15.844Z |  | FALSE
DI-34 | The device shall have an operating temperature between 0 and 40 degrees Celsius. - See DI17? | UN-12 |  |  | RSD Guardian 0.9x | 2024-08-13T15:05:41.646Z |  | FALSE
DI-35 | The device shall not incorporate external connections other than POE | UN-12 |  |  | RSD Guardian 0.9x | 2024-08-13T15:06:03.893Z |  | FALSE
DI-36 | The device shall incorporate a USB-C connection port to connect to the power supply. | UN-12 |  |  | RSD Guardian 0.9x | 2024-08-13T15:06:19.784Z |  | FALSE
DI-37 | The EM emissions from the device shall not exceed FCC standards. | UN-12 |  |  | RSD Guardian 0.9x | 2024-08-13T15:06:34.180Z |  | FALSE
DI-38 | The device shall be able to function under potential EM interference from other devices up to FCC Standards. | UN-12 |  |  | RSD Guardian 0.9x | 2024-08-13T15:06:48.069Z |  | FALSE
DI-39 | The firmware shall be able to automatically detect the following error states: Power Issues, Hardware Issues, SW Issues. | UN-13 |  |  | RSD Guardian 0.9x | 2024-08-13T15:07:06.251Z |  | FALSE
DI-40 | The firmware shall support internet connectivity. | UN-14 |  |  | RSD Guardian 0.9x | 2024-08-13T15:07:27.405Z |  | FALSE
DI-41 | The device shall be capable of firmware over the air updates | UN-14 |  |  | RSD Guardian 0.9x | 2024-08-13T15:07:45.637Z |  | FALSE
DI-42 | The firmware shall determine the patient's presence/absence and their relative movement from the signals received by the radar transmitter and the IR camera or the visual camera | UN-15 |  |  | RSD Guardian 0.9x | 2024-08-13T15:08:03.196Z |  | FALSE
DI-43 | The firmware shall be able to determine if a patient has fallen and send an alert to the cloud system | UN-16 |  |  | RSD Guardian 0.9x | 2024-08-13T15:09:34.177Z |  | FALSE
DI-44 | The cloud front end shall display a thermographic image of the patient | UN-17 |  |  | RSD Guardian 0.9x | 2024-08-13T15:10:07.458Z |  | FALSE
DI-45 | The Cloud front end shall display the last time a patient fell | UN-18 |  |  | RSD Guardian 0.9x | 2024-08-13T15:10:40.438Z |  | FALSE
DI-46 | The cloud front end shall display the last time a patient moved | UN-18 |  |  | RSD Guardian 0.9x | 2024-08-13T15:11:03.307Z |  | FALSE
DI-47 | The cloud front end shall have profiles for each patient with the ability to give a patient a profile image. Individual devices can be associated with patients | UN-19 |  |  | RSD Guardian 0.9x | 2024-08-13T15:13:01.194Z |  | FALSE
DI-48 | Patient profiles shall display the patient name, location group, profile image, and RSD Guardian Unit Number | UN-19 |  |  | RSD Guardian 0.9x | 2024-08-13T15:13:46.541Z |  | FALSE
DI-49 | Images can be added to patient profiles via drag and drop. | UN-19 |  |  | RSD Guardian 0.9x | 2024-08-13T15:14:03.328Z |  | FALSE
DI-50 | The cloud front end shall allow users to organize patients into groups. | UN-20 |  |  | RSD Guardian 0.9x | 2024-08-13T15:14:26.037Z |  | FALSE
DI-51 | The cloud front end shall allow users to enable or disable to devices, view device issues, and edit associated information such as the name of the patient being monitored | UN-21 |  |  | RSD Guardian 0.9x | 2024-08-13T15:15:21.510Z |  | FALSE
DI-52 | Device shall incorporate a visual camera?? | UN-1, UN-3
DI-53 | The motion and fall detection capabilities will cover the entire 10x10ft room from an elevation of 10 ft, which is approximately a distance of at least 13 feet | UN-22
DI-54 | The motion and fall detection system will have less than 1 false alarm per 12 hours using test data (I made this up) | UN-23
changed or added
do we need this Design Input?
delete
Note on C1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: MY Addition, not exported to CSV